## QA-QC Data
QA-QC Results
Calibration Curve Slope Precision (± 2s)
| Br | Cl | SO4
| 1.3 | 1.3 | 1.5
Blank Verification (mg L-1)
| Br | Cl | SO4
Control Limit | 0.096 | 2.9 | 0.048
Initial Calibration Blank | non-detect | 0.01 | non-detect
Continuing Calibration Blank (1) | non-detect | 0.04 | 0.028
Continuing Calibration Blank (2) | non-detect | 0.18 | non-detect
Continuing Calibration Blank (3) | non-detect | 0.16 | 0.006
Initial Calibration Verification (mg L-1)
| Br | Cl | SO4
Actual Concentration | 1.98 | 60.3 | 1.16
Measured Concentration | 1.86 | 60.6 | 1.09
Percent Deviation | -6.4516 | 0.49505 | -6.422
Continued Calibration Verification (mg L-1)
| Br | Cl | SO4
CCV#1
Actual Concentration | 1.9 | 57 | 0.953
Measured Concentration | 1.8 | 58.4 | 0.905
Percent Deviation | -5.5556 | 2.3973 | -5.3039
CCV#2
Actual Concentration | 1.9 | 57 | 0.953
Measured Concentration | 1.85 | 60.4 | 0.938
Percent Deviation | -2.7027 | 5.6291 | -1.5991
Replicates and Duplicates - Relative Percent Differences
| Br | Cl | SO4
Lab Replicates
Masterson.UT.20240401.001 - CCU 12-1 | 0.072 | 0.19 | 0.17
Masterson.UT.20240401.002 - CCU 17-1 | 4.4 | 3.7 | 4.1
Masterson.UT.20240401.003 - CCU 28-2 | 2.7 | 2.5 | 2.3
Masterson.UT.20240401.004 - KSF 10-1 | 1.3 | 1.2 | 0.3
Masterson.UT.20240401.005 - TM 43-18H | 2.8 | 2.5 | 0.12
Masterson.UT.20240401.006 - S 16-2 | 4.4 | 3.8 | 8 | s
Masterson.UT.20240401.007 - F 28-11 | 5.7 | 5 | 5.3
Masterson.UT.20240401.008 - UT 34-13J-D3 | 2.6 | 0.61 | <PQL
Masterson.UT.20240401.009 - UT 13H-16-55 | 5.5 | 2.2 | 30
Laboratory Duplicate
Masterson.UT.20240401.001 - CCU 12-1 | 3060 | 169000 | 64.6
LAB DUP Masterson.UT.20240401.001 - CCU 12-1 | 3100 | 171000 | 65.8
Percent Deviation | 1.2903 | 1.1696 | 1.8237
Reference Standards (mg L-1)
| Br | Cl | SO4
Multi Analyte IC Solution
Actual Concentration | 20 | 20 | 30
Measured Concentration | 20.7 | 20.6 | 32.2
Percent Deviation | 3.3816 | 2.9126 | 6.8323
M234
Actual Concentration | 0.126 | 34.9 | 8.59
Measured Concentration | 0.132 | 36.7 | 9.19
Percent Deviation | 4.5455 | 4.9046 | 6.5288
M238
Actual Concentration | 0.614 | 87.7 | 14.1
Measured Concentration | 0.678 | 89.8 | 15.1
Percent Deviation | 9.4395 | 2.3385 | 6.6225
M242
Actual Concentration | 0.511 | 41.9 | 8.41
Measured Concentration | 0.507 | 42.9 | 9.03
Percent Deviation | -0.78895 | 2.331 | 6.866
M248
Actual Concentration | 0.424 | 128 | 79.2
Measured Concentration | 0.466 | 130 | 83.4
Percent Deviation | 9.0129 | 1.5385 | 5.036
Matrix Spikes (mg L-1)
| Br | Cl | SO4
MS - DUP Masterson.UT.20240401.001 - CCU 12-1
1Measured DUP Masterson.UT.20240401.001 - CCU 12-1 (mg L-1) | 7.25 | 401 | 0.154
Spike Concentration (mg L-1) | 2 | 60 | 1
Theoretical MS - DUP Masterson.UT.20240401.001 - CCU 12-1 (mg L-1) | 9.25 | 461 | 1.154
Measured MS - DUP Masterson.UT.20240401.001 - CCU 12-1 (mg L-1) | 9.58 | 475 | 1.11
Percent Recovery | 103.5676 | 103.0369 | 96.1872
1. Corresponds to 90% (volume basis) measured values post dilution
Notes
A. Yellow highlight indicates value outside of acceptance criteria.
B. QA-QC Acceptance Criteria: | Notes
1. Calibration curve slope precise to ±2% at the 2s level.
2. Blanks <50% of PQL.
3. ICV values within ±10% of true value.
4. CCV values within ±10% of true value.
5. Lab replicate RPD <±15%. | Sample 'UT 13H-16-55' has a relative percent difference of 30%, and concentrations on the lower end of the calibration curve have acceptance criteria of 50%, which this sample falls within.
6. Lab duplicates have no acceptance criteria.
7. Reference standards within ±15% of true value.
8. Matrix spike recovery within 80% - 120%.
C. Green highlight indicates concentration < PQL, therefore data is semi-quantitative.
D. Blue highlight indicates concentration > highest calibration standard, therefore data is semi-quantitative.
Compound | Calibration Range (mg L-1)
Br | 0.2 - 20
Cl | 6 - 600
SO4 | 0.1 - 10

## Data Summary
BRInE ID | Submitter ID | Bromide (mg L-1) | Bromide ± 2s | Chloride (mg L-1) | Chloride ± 2s | Sulfate (mg L-1) | Sulfate ± 2s
Masterson.UT.20240401.001 | CCU 12-1 | 3060 | 110 | 169000 | 17000 | 64.6 | 2.1
Masterson.UT.20240401.002 | CCU 17-1 | 1900 | 170 | 106000 | 13000 | 60 | 5
Masterson.UT.20240401.003 | CCU 28-2 | 989 | 61 | 58700 | 6500 | 62.3 | 3.2
Masterson.UT.20240401.004 | KSF 10-1 | 2830 | 120 | 194000 | 20000 | 127 | 4
Masterson.UT.20240401.005 | TM 43-18H | 2200 | 140 | 149000 | 17000 | 43.5 | 1.6
Masterson.UT.20240401.006 | S 16-2 | 1920 | 170 | 204000 | 26000 | 165 | 25
Masterson.UT.20240401.007 | F 28-11 | 988 | 107 | 119000 | 17000 | 165 | 17
Masterson.UT.20240401.008 | UT 34-13J-D3 | 28.4 | 1.5 | 4740 | 470 | <1.15 | <PQL
Masterson.UT.20240401.009 | UT 13H-16-55 | 18 | 1.9 | 2810 | 290 | 2.91 | 1.67
Flags
1. Green highlight indicates concentration < PQL, therefore data is semi-quantitative.
2. Blue highlight indicates concentration > highest calibration standard, therefore data is semi-quantitative.
3. Standard Operating Procedure requires samples to be unacidified and filtered. Only two unacidified samples were available ('UT 34-13J-D3' and 'UT 13H-16-55'). All other samples are filtered and acidified.
Notes
A. (i) Sample 'UT 13H-16-55' has a relative percent difference of 30% for Sulfate, and concentrations on the lower end of the calibration curve have acceptance criteria of 50%, which this sample falls within. Sample was also cloudy in appearance, maybe soapy, refiltered in lab.
(ii) Chromatograms show additional ions present.
B. Error limits correspond to the ±2𝛔 precision uncertainties.
C. Typical elution times
| Compound | Elution time (min)
| Br | 6.2
| Cl | 3.7
| SO4 | 8.9
D. Calibration standard concentrations and PQLs
| Compound | PQL (mg L-1) | Calibration Range (mg L-1)
| Br | 0.19 | 0.2 - 20
| Cl | 5.8 | 6 - 600
| SO4 | 0.097 | 0.1 - 10
E. Sample dilution factors. Stated concentrations already reflect dilution correction.
| Brine ID | Submitter ID | Dilution Factor
| Masterson.UT.20240401.001 | CCU 12-1 | 357
| Masterson.UT.20240401.002 | CCU 17-1 | 208
| Masterson.UT.20240401.003 | CCU 28-2 | 119
| Masterson.UT.20240401.004 | KSF 10-1 | 417
| Masterson.UT.20240401.005 | TM 43-18H | 313
| Masterson.UT.20240401.006 | S 16-2 | 417
| Masterson.UT.20240401.007 | F 28-11 | 250
| Masterson.UT.20240401.008 | UT 34-13J-D3 | 11.9
| Masterson.UT.20240401.009 | UT 13H-16-55 | 9.26

## Traceability
Analyst: | A. Herzberg
Date Analyzed: | 2024-04-09
Secondary Review by: | J. Chenault
SOP: | ERP-BRInE-SOP 12
Version: | 06
Balance ID: | Brine-balance-B146459013, Brine-balance- B15055081
Balance Calibration Date: | February 28th, 2024
Pipette ID: | p.10-100.001; p.100-1000.001; p.500-5000.001
Pipette Calibration Date: | March 21st, 2024
Instrument Conditions
Column | ThermoFisher IonPac AS18- Fast 2x150mm
Guard Column | ThermoFisher IonPac AG18- Fast 2x30mm
KOH Eluent (mM) | 16
Column Temp | 35
Cell Temp | 35
Run Time (min) | 16
Flow Rate | 0.25 mL/min isocratic
Detection | AERS_2mm, 5.0 Hz, Suppressor 13 mA
System Backpressure | ~2500 psi
Background conductance | > 0.2 and <1.0
Notes: