## QA-QC Data
QA-QC Results
Calibration Curve Slope Precision (± 2s)
| B | Ba | Ca | Fe | K | Li | Mg | Mn | Na | Si | Sr
| 1.1 | 2 | 1.2 | 0.8 | 2 | 1.4 | 0.7 | 0.9 | 1.5 | 1.3 | 2.7
Blank Verification (mg L-1)
| B | Ba | Ca | Fe | K | Li | Mg | Mn | Na | Si | Sr
Control limit | 0.005 | 0.006 | 0.504 | 0.007 | 0.09 | 0.005 | 0.06 | 0.004 | 0.798 | 0.005 | 0.03
Lab Blank #1 | non-detect | non-detect | 0.01 | non-detect | 0.128 | non-detect | 0.002 | non-detect | non-detect | 0.003 | non-detect
Lab Blank #2 | 0.001 | non-detect | non-detect | non-detect | 0.063 | 0.004 | 0.002 | non-detect | non-detect | 0.001 | non-detect
Lab Blank #3 | non-detect | non-detect | 0.01 | non-detect | non-detect | non-detect | 0.003 | non-detect | non-detect | non-detect | non-detect
Lab Blank #4 | non-detect | non-detect | non-detect | non-detect | non-detect | 0.003 | non-detect | non-detect | 0.01 | non-detect | non-detect
Lab Blank #5 | non-detect | non-detect | non-detect | non-detect | 0.067 | non-detect | 0.004 | non-detect | non-detect | 0.002 | non-detect
Lab Blank #6 | non-detect | non-detect | non-detect | non-detect | 0.099 | non-detect | non-detect | non-detect | non-detect | non-detect | non-detect
Lab Blank #7 | non-detect | non-detect | 0.01 | non-detect | non-detect | non-detect | non-detect | non-detect | non-detect | non-detect | non-detect
Initial Calibration Verification (mg L-1)
| B | Ba | Ca | Fe | K | Li | Mg | Mn | Na | Si | Sr
Actual Concentration | 1.25 | 1.5 | 126 | 1.75 | 22.6 | 1.25 | 15.1 | 1 | 200 | 1.25 | 7.51
Measured Concentration | 1.25 | 1.51 | 129 | 1.79 | 22.4 | 1.21 | 15.3 | 1.02 | 202 | 1.26 | 7.64
Percent Deviation | 0 | 0.66225 | 2.3256 | 2.2346 | -0.89286 | -3.3058 | 1.3072 | 1.9608 | 0.9901 | 0.79365 | 1.7016
Continued Calibration Verification (mg L-1)
| B | Ba | Ca | Fe | K | Li | Mg | Mn | Na | Si | Sr
CCV#1
Actual Concentration | 0.502 | 0.601 | 50.4 | 0.702 | 9.04 | 0.501 | 6.02 | 0.4 | 79.8 | 0.5 | 3
Measured Concentration | 0.488 | 0.608 | 51.6 | 0.726 | 9.18 | 0.51 | 6.12 | 0.404 | 81.1 | 0.498 | 2.99
Percent Deviation | -2.8689 | 1.1513 | 2.3256 | 3.3058 | 1.5251 | 1.7647 | 1.634 | 0.9901 | 1.603 | -0.40161 | -0.33445
CCV#2
Actual Concentration | 0.502 | 0.601 | 50.4 | 0.702 | 9.04 | 0.501 | 6.02 | 0.4 | 79.8 | 0.5 | 3
Measured Concentration | 0.498 | 0.618 | 50.3 | 0.715 | 9.07 | 0.501 | 5.96 | 0.41 | 80.1 | 0.498 | 2.92
Percent Deviation | -0.80321 | 2.7508 | -0.19881 | 1.8182 | 0.33076 | 0 | -1.0067 | 2.439 | 0.37453 | -0.40161 | -2.7397
Duplicates and Replicates (mg L-1)
| B | Ba | Ca | Fe | K | Li | Mg | Mn | Na | Si | Sr
Laboratory Duplicates
Masterson.UT.20240401.001 - CCU 12-1 | 44.9 | 4.2 | 64800 | 63.4 | 1170 | 19.9 | 7770 | 10.6 | 32000 | 2.33 | 2160
DUPILCATE Masterson.UT.20240401.001 - CCU 12-1 | 29 | 4.67 | 71400 | 69.6 | 1280 | 22.2 | 8530 | 11.6 | 35600 | 2.34 | 2370
Percent Deviation | -54.8276 | 10.0642 | 9.2437 | 8.908 | 8.5938 | 10.3604 | 8.9097 | 8.6207 | 10.1124 | 0.42735 | 8.8608
Laboratory Replicates
Masterson.UT.20240401.005 - TM 43-18H | 26.6 | 11.6 | 63900 | 46.3 | 4440 | 24.3 | 5960 | 16.7 | 13200 | 0.952 | 2780
REPLICATE Masterson.UT.20240401.005 - TM 43-18H | 26.8 | 11.9 | 65700 | 47.3 | 4480 | 24.2 | 6200 | 17.2 | 13100 | 1.03 | 2770
Percent Deviation | 0.74627 | 2.521 | 2.7397 | 2.1142 | 0.89286 | -0.41322 | 3.871 | 2.907 | -0.76336 | 7.5728 | -0.36101
Reference Standards (mg L-1)
| B | Ba | Ca | Fe | K | Li | Mg | Mn | Na | Si | Sr
Certified Seawater
Actual Concentration | 5.2 | - | 417 | - | 396 | - | - | - | 10900 | 4.2 | 12.5
Measured Concentration | 5.3 | - | 408 | - | 297 | - | - | - | 10800 | 4.01 | 12
Percent Deviation | B: 1.8868; Ca: -2.2059; K: -33.3333; Na: -0.92593; Si: -4.7382; Sr: -4.1667
M234
Actual Concentration | - | - | 15.7 | - | 1.81 | - | 3.6 | - | 18 | - | -
Measured Concentration | - | - | 16.2 | - | 1.72 | - | 3.78 | - | 18.6 | - | -
Percent Deviation | Ca: 3.0864; K: -5.2326; Mg: 4.7619; Na: 3.2258
M238
Actual Concentration | 0.15 | - | 40.9 | - | 5.99 | - | 10.1 | - | 22.3 | - | -
Measured Concentration | 0.142 | - | 31.6 | - | 5.56 | - | 10.1 | - | 22.8 | - | -
Percent Deviation | B: -5.6338; Ca: -29.4304; K: -7.7338; Mg: 0; Na: 2.193
T241
Actual Concentration | - | - | 18.5 | 0.232 | 1.15 | - | 6.76 | - | 16 | - | -
Measured Concentration | - | - | 19.1 | 0.239 | <0.9 | - | 7.03 | - | 16.2 | - | -
Percent Deviation | Ca: 3.1414; Fe: 2.9289; Mg: 3.8407; Na: 1.2346
T249
Actual Concentration | 0.211 | 0.11 | 83.4 | 0.0795 | 5.32 | - | 19.8 | 0.226 | 69.7 | - | 0.404
Measured Concentration | 0.202 | 0.112 | 83.6 | 0.0796 | 3.76 | - | 21.1 | 0.232 | 71.7 | - | 0.404
Percent Deviation | -4.4554 | 1.7857 | 0.23923 | 0.12563 | -41.4894 |  | 6.1611 | 2.5862 | 2.7894 |  | 0
Matrix Spikes (mg L-1)
| B | Ba | Ca | Fe | K | Li | Mg | Mn | Na | Si | Sr
MS - DUP Masterson.UT.20240401.001 - CCU 12-1
1Measured DUP Masterson.UT.20240401.001 - CCU 12-1 (mg L-1) | 0.078 | 0.013 | 193 | 0.188 | 3.46 | 0.06 | 23 | 0.031 | 96.3 | 0.006 | 6.4
Spike Concentration (mg L-1) | 0.251 | 0.3 | 25.2 | 0.351 | 4.52 | 0.251 | 3.01 | 0.2 | 39.9 | 0.25 | 1.5
Theoretical MS DUP Masterson.UT.20240401.001 - CCU 12-1 (mg L-1) | 0.329 | 0.313 | 218.2 | 0.539 | 7.98 | 0.311 | 26.01 | 0.231 | 136.2 | 0.256 | 7.9
Measured MS DUP Masterson.UT.20240401.001 - CCU 12-1 (mg L-1) | 0.364 | 0.309 | 212 | 0.524 | 7.95 | 0.314 | 25.6 | 0.225 | 134 | 0.25 | 7.93
Percent Recovery | 110.6383 | 98.722 | 97.1586 | 97.2171 | 99.6241 | 100.9646 | 98.4237 | 97.4026 | 98.3847 | 97.6562 | 100.3797
1. Corresponds to 90% (volume basis) measured values post dilution
Notes
A. Yellow highlight indicates value outside of acceptance criteria.
B. QA-QC Acceptance Criteria: | B: Exceeds Acceptance Criteria
1. Calibration curve slope precise to ±2% at the 2s level | B: Strontium slope precision of 2.7; data is flagged.
2. Blanks <10% of PQL | B: Potassium for laboratory blanks 1 and 6.
3. ICV values within ±10% of true value.
4. CCV values within ±10% of true value.
5. Laboratory duplicates have no acceptance criteria.
6. Laboratory replicates within ±20% of each other.
7. Reference standards within ±15% of true value. | B: Calcium for M238 was underestimated. Potassium for Seawater and T249 was underestimated.
8. Matrix spike recovery within 70% - 130%.
C. Green highlight indicates concentration < PQL, therefore data is semi-quantitative.
D. Blue highlight indicates concentration > highest calibration standard, therefore data is semi-quantitative.
Element | B: Calibration Range (mg L-1)
B | B: 0.05 - 2.5
Ba | B: 0.06 - 3
Ca | B: 5.0 - 250
Fe | B: 0.07 - 3.5
K | B: 0.9 - 45
Li | B: 0.05 - 2.5
Mg | B: 0.6 - 30
Mn | B: 0.04 - 2.0
Na | B: 8 - 400
Si | B: 0.05 - 2.5
Sr | B: 0.3 - 15

## Data Summary
BRInE ID | Sample ID | [B] (mg L-1) | ± 2s | [Ba] (mg L-1) | ± 2s | [Ca] (mg L-1) | ± 2s | [Fe] (mg L-1) | ± 2s | [K] (mg L-1) | ± 2s | [Li] (mg L-1) | ± 2s | [Mg] (mg L-1) | ± 2s | [Mn] (mg L-1) | ± 2s | [Na] (mg L-1) | ± 2s | [Si] (mg L-1) | ± 2s | 3[Sr] (mg L-1) | ± 2s
Masterson.UT.20240401.001 | CCU 12-1 | 44.9 | 1.2 | <18.7 | <PQL | 64800 | 5500 | 63.4 | 1.2 | 1170 | 150 | 19.9 | 2.7 | 7770 | 240 | <12.5 | <PQL | 32000 | 3300 | <15.6 | <PQL | 2160 | 170
Masterson.UT.20240401.002 | CCU 17-1 | 17.7 | 0.9 | <10.8 | <PQL | 35300 | 3200 | 30.6 | 1 | 696 | 125 | 12.1 | 1.2 | 4410 | 210 | <7.16 | <PQL | 18000 | 2000 | <8.95 | <PQL | 1160 | 90
Masterson.UT.20240401.003 | CCU 28-2 | 9.47 | 0.36 | 9.09 | 0.83 | 19600 | 1800 | 7.37 | 0.32 | 495 | 41 | 7.25 | 0.51 | 2190 | 60 | <4.16 | <PQL | 11200 | 1400 | <5.20 | <PQL | 651 | 52
Masterson.UT.20240401.004 | KSF 10-1 | 22.2 | 0.9 | <21.4 | <PQL | 55600 | 4500 | 97.5 | 4.8 | 1460 | 90 | <17.9 | <PQL | 6890 | 350 | <14.3 | <PQL | 45700 | 5100 | <17.9 | <PQL | 1970 | 150
Masterson.UT.20240401.005 | TM 43-18H | 26.6 | 1.3 | <15.8 | <PQL | 63900 | 6800 | 46.3 | 1.1 | 4440 | 390 | 24.3 | 1.8 | 5960 | 180 | 16.7 | 0.7 | 13200 | 1300 | <13.2 | <PQL | 2780 | 310
Masterson.UT.20240401.006 | S 16-2 | 54.5 | 4 | <21.4 | <PQL | 47900 | 3800 | 95.6 | 2.7 | 2710 | 240 | 48.5 | 4.6 | 4530 | 190 | <14.3 | <PQL | 73600 | 8200 | <17.9 | <PQL | 2310 | 200
Masterson.UT.20240401.007 | F 28-11 | 61.8 | 3.4 | <12.5 | <PQL | 24400 | 2000 | 222 | 7 | 2630 | 310 | 35.9 | 2.3 | 2980 | 120 | 16.9 | 0.6 | 37400 | 4200 | <10.4 | <PQL | 1080 | 100
Masterson.UT.20240401.008 | UT 34-13J-D3 | 5.95 | 0.34 | 9.09 | 0.37 | <50.4 | <PQL | <0.702 | <PQL | 29.7 | 1.9 | 5.25 | 0.25 | <6.02 | <PQL | <0.400 | <PQL | 3850 | 460 | 13.7 | 0.5 | 5.65 | 0.44
Masterson.UT.20240401.009 | UT 13H-16-55 | 10.2 | 0.2 | 24.2 | 1.5 | <40.6 | <PQL | 2.17 | 0.06 | 61.3 | 6.6 | 6.41 | 0.09 | <4.86 | <PQL | <0.322 | <PQL | 2210 | 270 | 17.9 | 0.8 | 3.05 | 0.17
Flags
1. Green highlight indicates concentration < PQL, therefore data is semi-quantitative.
2. Blue highlight indicates concentration > highest calibration standard, therefore data is semi-quantitative.
3. Strontium's 2.7% slope precision exceeds acceptance criteria; the data is flagged with this note. All other QC is within the acceptable range.
Notes
A. (i) Potassium QC that exceeds acceptance criteria are laboratory blanks 1 and 6, and reference standards Seawater and T249 is underestimated. The blanks are below PQL. Other QCs within acceptable range.
(ii) Calcium QC that exceeds acceptance criteria for reference standard M238 is underestimated. Other Qc within acceptable range.
(iii) Sample 'UT 13H-16-55' was cloudy in appearance, maybe soapy, refiltered in lab prior to analysis.
B. Error limits correspond to the ±2𝛔 precision uncertainties.
C. Element concentrations based on:
| Element | Line (nm)
| B | 249.773
| Ba | 455.403
| Ca | 315.887
| Fe | 259.94
| Li | 670.791
| Mg | 279.079
| Mn | 257.61
| K | 766.49
| Na | 588.995
| Si | 251.611
| Sr | 407.771
D. Calibration standard concentrations and PQLs
| Element | Calibration Range (mg L-1)
| B | 0.05 - 2.5
| Ba | 0.06 - 3
| Ca | 5.0 - 250
| Fe | 0.07 - 3.5
| K | 0.9 - 45
| Li | 0.05 - 2.5
| Mg | 0.6 - 30
| Mn | 0.04 - 2.0
| Na | 8 - 400
| Si | 0.05 - 2.5
| Sr | 0.3 - 15
E. Sample dilution factors. Stated concentrations already reflect dilution correction.
| Brine ID | Submitter ID | Dilution Factor
| Masterson.UT.20240401.001 | CCU 12-1 | 313
| Masterson.UT.20240401.002 | CCU 17-1 | 179
| Masterson.UT.20240401.003 | CCU 28-2 | 104
| Masterson.UT.20240401.004 | KSF 10-1 | 357
| Masterson.UT.20240401.005 | TM 43-18H | 263
| Masterson.UT.20240401.006 | S 16-2 | 357
| Masterson.UT.20240401.007 | F 28-11 | 208
| Masterson.UT.20240401.008 | UT 34-13J-D3 | 10
| Masterson.UT.20240401.009 | UT 13H-16-55 | 8.1

## Traceability
Analyst: | A. Herzberg
Date samples run: | 2024-04-10
Secondary Review by: | J. Chenault
Instrument: | Horiba Ultima Expert ICP-OES
SOP: | ERP-BRInE-SOP 13
Version: | 02
Balance ID: | Brine-balance-B146459013, Brine-balance- B15055081
Balance calibration date: | February 28th, 2024
Pipette ID: | p.10-100.001; p.100-1000.001; p.500-5000.001
Pipette calibration date: | March 21st, 2024
Instrument Conditions
RF Power (W) | 1250
Pump Speed (rpm) | 31
Ar Plasma Gas Flow (l/min) | 14.5
Ar Aux Gas Flow (l/min) | 0.6
Ar Sheath Gas Flow (l/min) | 0.3
N2 purge flow (l/min) | off
Nebulizer Flow (ml/min) | 0.84
Nebulizer Pressure (psi) | 1.94
Notes: