## QA-QC Report
QA-QC Results
Calibration Curve Slope Precision (± 2s)
| Br | Cl | SO4
| 0.9 | 0.9 | 1.1
Blank Verification (mg L-1)
| Br | Cl | SO4
Control Limit | 0.1 | 2.5 | 0.38
Initial Calibration Blank | <0.006 | <0.006 | <0.009
Continuing Calibration Blank (1) | <0.006 | 0.02 | 0.018
Continuing Calibration Blank (2) | <0.006 | 0.03 | <0.009
Continuing Calibration Blank (3) | <0.006 | <0.006 | <0.009
Initial Calibration Verification (mg L-1)
| Br | Cl | SO4
Actual Concentration | 2 | 50 | 7
Measured Concentration | 1.92 | 49.4 | 7.04
Percent Deviation | -4.1667 | -1.2146 | 0.56818
Continued Calibration Verification (mg L-1)
| Br | Cl | SO4
CCV#1
Actual Concentration | 2 | 50 | 7.61
Measured Concentration | 1.89 | 48.9 | 7.52
Percent Deviation | -5.8201 | -2.2495 | -1.1968
CCV#2
Actual Concentration | 2 | 50 | 7.61
Measured Concentration | 1.86 | 48.3 | 7.4
Percent Deviation | -7.5269 | -3.5197 | -2.8378
Replicates and Duplicates - Relative Percent Differences
| Br | Cl | SO4
Laboratory Replicates
Masterson.UT.20240725.001: FD43 | 0.4 | 0.3 | 0.5
Masterson.UT.20240725.002: FD50 | 3.2 | 0.3 | 0.5
Masterson.UT.20240725.003: FD51 | 1.9 | 0.6 | 0.7
Masterson.UT.20240725.004: FD68 | 0.1 | 0.4 | 0.3
Masterson.UT.20240725.005: FD75 | 1 | 1.6 | 1.1
Masterson.UT.20240725.006: FD70 | 0.5 | 2.2 | 2.1
Masterson.UT.20240725.007: FD86 | 3.3 | 0.8 | 3.6
Masterson.UT.20240725.008: FD88 | 0.1 | 6.9 | 4.7
Laboratory Duplicate
Masterson.UT.20240725.001: FD43 | 15.1 | 2260 | 82.7
DUP Masterson.UT.20240725.001: FD43 | 14.1 | 2120 | 77.9
Percent Deviation | -7.0922 | -6.6038 | -6.1617
Reference Standards (mg L-1)
| Br | Cl | SO4
M238
Actual Concentration | 0.614 | 87.7 | 14.1
Measured Concentration | 0.639 | 87.8 | 14.6
Percent Deviation | 3.9124 | 0.1139 | 3.4247
M242
Actual Concentration | 0.511 | 41.9 | 8.41
Measured Concentration | 0.511 | 40.7 | 8.51
Percent Deviation | 0 | -2.9484 | 1.1751
M248
Actual Concentration | 0.424 | 128 | 79.2
Measured Concentration | 0.415 | 124 | 74.9
Percent Deviation | -2.1687 | -3.2258 | -5.741
Multi Analyte IC Solution
Actual Concentration | 20 | 20 | 30
Measured Concentration | 19.1 | 19.7 | 30
Percent Deviation | -4.712 | -1.5228 | 0
Matrix Spikes (mg L-1)
| Br | Cl | SO4
MS: Masterson.UT.20240725.001: FD43 | 3.47 | 280 | 16.1
1LD: Masterson.UT.20240725.001: FD43 | 1.53 | 229 | 8.41
Spike Added | 2 | 50 | 7.61
Percent Recovery | 97 | 102 | 101.0512
1. Corresponds to 90% (volume basis) measured values post dilution
Notes
Analyte | Calibration Range (mg L-1)
Br | 0.2 - 20
Cl | 5 - 500
SO4 | 0.8 - 76
A. The green highlight indicates the concentrations are less than the PQL, the lowest calibration standard and the blue highlight indicates the concentrations are greater than the highest calibration standard. Non-detect indicates that the analyte is not detected with this method.
B. Yellow highlight indicates value outside of acceptance criteria defined below.
QA-QC Acceptance Criteria: | Notes
1. Slope precision ≤2% at the 2s level. | N/A
2. Blanks <50% of PQL. | N/A
3. ICV values within ±10% of true value. | N/A
4. CCV values within ±10% of true value. | N/A
5. Lab replicate RPD <±15%. | N/A
6. Lab duplicates have no acceptance criteria. | N/A
7. Reference standards within ±15% of true value. | N/A
8. Matrix spike recovery within 70-120% | N/A

## Data Summary
|  | Br | ± Br Error | Cl | ± Cl Error | SO42- | ± SO42- Error
BRInE_ID | Submitter_ID | mg/L | mg/L | mg/L | mg/L | mg/L | mg/L
Masterson.UT.20240725.001 | FD43 | 15.1 | 0.3 | 2260 | 150 | 82.7 | 3.2
Masterson.UT.20240725.002 | FD50 | 24.3 | 1.5 | 3320 | 220 | 247 | 11
Masterson.UT.20240725.003 | FD51 | 22.1 | 0.9 | 2760 | 190 | 365 | 18
Masterson.UT.20240725.004 | FD68 | 26.4 | 0.5 | 3380 | 220 | 311 | 14
Masterson.UT.20240725.005 | FD75 | 34.3 | 0.9 | 5060 | 380 | 79.3 | 3.2
Masterson.UT.20240725.006 | FD70 | 27.9 | 0.5 | 3980 | 320 | 222 | 13
Masterson.UT.20240725.007 | FD86 | 62.1 | 4.1 | 10100 | 700 | 38.3 | 2.9
Masterson.UT.20240725.008 | FD88 | 63.5 | 1.1 | 10400 | 1600 | 34.1 | 3.2
Flags
N/A
Notes
Concentrations below the method detection limit (MDL), below the lowest calibration standard (PQL), or above the highest standard (Standard) are reported as the concentration of the MDL, PQL, or high standard with the sample's dilution factor.
Error limits correspond to the ± 2𝛔 precision uncertainties.
A. QA-QC and Sample Considerations
N/A
B. Average elution times
Analyte | Elution time (min)
Br | 3.6
Cl | 5.8
SO42- | 8.1
C. Calibration standards range and MDL concentrations
Analyte | MDL (mg L-1) | Calibration Range (mg L-1)
Br | 0.006 | 0.2 - 20
Cl | 0.006 | 5 - 500
SO42- | 0.009 | 0.8 - 76
D. Sample dilution factors. Stated concentrations already reflect dilution correction.
Brine ID | Submitter ID | Dilution Factor
Masterson.UT.20240725.001 | FD43 | 8.3
Masterson.UT.20240725.002 | FD50 | 10
Masterson.UT.20240725.003 | FD51 | 8.3
Masterson.UT.20240725.004 | FD68 | 12.5
Masterson.UT.20240725.005 | FD75 | 12.5
Masterson.UT.20240725.006 | FD70 | 12.5
Masterson.UT.20240725.007 | FD86 | 25
Masterson.UT.20240725.008 | FD88 | 25

## Traceability
ERP-BRInE-SOP 12
SOP Version | 06
BRInE Sample ID | Masterson.UT.20240725.001-008
Sequence Name | Masterson.UT.20240725.001-008-R2
Date of Analysis | 8/9/2024
Analyst | A. Herzberg
Reviewer | J. Chenault
Method Conditions
Eluent (mM KOH) | 16
Suppressor (mA) | 13
Column Temp ( °C) | 35
Cell Temp ( °C) | 35
Flow rate (mL/min) | 0.25
Run Time (min) | 16
Notes
N/A