## QA-QC Data
QA-QC Results
Calibration Curve Slope Precision (± 2s)
| B | Ba | Ca | Fe | K | Li | Mg | Mn | Na | Si | Sr
| 1.5 | 1.7 | 1.3 | 1.0 | 1.7 | 1.7 | 1.3 | 0.8 | 0.9 | 0.8 | 1.3
Blank Verification (mg L-1)
| B | Ba | Ca | Fe | K | Li | Mg | Mn | Na | Si | Sr
Control limit | 0.005 | 0.002 | 0.1 | 0.002 | 0.09 | 0.005 | 0.02 | 0.002 | 0.9 | 0.03 | 0.03
Lab Blank #1 | 0.001 | <MDL | <MDL | <MDL | <MDL | <MDL | 0.0070658 | <MDL | <MDL | <MDL | <MDL
Lab Blank #2 | <MDL | <MDL | <MDL | <MDL | <MDL | <MDL | <MDL | <MDL | <MDL | <MDL | <MDL
Lab Blank #3 | <MDL | <MDL | 0.005 | <MDL | <MDL | <MDL | <MDL | <MDL | <MDL | <MDL | 0.00080032
Lab Blank #4 | <MDL | <MDL | <MDL | <MDL | 0.13422 | <MDL | <MDL | <MDL | <MDL | <MDL | <MDL
Lab Blank #5 | 0.001 | <MDL | <MDL | <MDL | 0.081275 | <MDL | <MDL | <MDL | <MDL | <MDL | <MDL
Lab Blank #6 | <MDL | 0.0008 | 0.007 | <MDL | <MDL | <MDL | 0.012963 | <MDL | <MDL | <MDL | 0.001
Lab Blank #7 | <MDL | <MDL | 0.001 | <MDL | <MDL | 0.00070562 | <MDL | <MDL | <MDL | <MDL | <MDL
Initial Calibration Verification (mg L-1)
| B | Ba | Ca | Fe | K | Li | Mg | Mn | Na | Si | Sr
Actual Concentration | 1.25 | 0.5 | 25 | 0.5 | 22.5 | 1.25 | 5 | 0.5 | 225 | 7.5 | 7.5
Measured Concentration | 1.26 | 0.505 | 24.1 | 0.5 | 22.1 | 1.26 | 5.01 | 0.505 | 227 | 7.37 | 7.52
Percent Deviation | 0.79365 | 0.9901 | -3.7344 | 0 | -1.81 | 0.79365 | 0.1996 | 0.9901 | 0.88106 | -1.7639 | 0.26596
Continued Calibration Verification (mg L-1)
| B | Ba | Ca | Fe | K | Li | Mg | Mn | Na | Si | Sr
CCV#1
Actual Concentration | 0.5 | 0.2 | 10 | 0.2 | 9 | 0.5 | 2 | 0.2 | 90 | 3 | 3
Measured Concentration | 0.486 | 0.2 | 9.27 | 0.196 | 9.01 | 0.503 | 1.96 | 0.194 | 92.6 | 2.9 | 2.88
Percent Deviation | -2.8807 | 0 | -7.8749 | -2.0408 | 0.11099 | 0.59642 | -2.0408 | -3.0928 | 2.8078 | -3.4483 | -4.1667
CCV#2
Actual Concentration | 0.5 | 0.2 | 10 | 0.2 | 9 | 0.5 | 2 | 0.2 | 90 | 3 | 3
Measured Concentration | 0.485 | 0.203 | 9.33 | 0.2 | 9.02 | 0.505 | 2 | 0.196 | 93.8 | 2.97 | 2.92
Percent Deviation | -3.0928 | 1.4778 | -7.1811 | 0 | 0.22173 | 0.9901 | 0 | -2.0408 | 4.0512 | -1.0101 | -2.7397
Duplicates and Replicates (mg L-1)
| B | Ba | Ca | Fe | K | Li | Mg | Mn | Na | Si | Sr
Laboratory Duplicates
Masterson.UT.20240725.001: FD43 | 11.3 | 1.14 | 23.9 | 0.328 | 12.1 | 1.66 | 3.8 | 0.0196 | 2030 | 39.8 | 4.38
DUP Masterson.UT.20240725.001: DUP FD43 | 11.3 | 1.15 | 23.2 | 0.328 | 11.5 | 1.7 | 3.72 | 0.0197 | 1990 | 39.5 | 4.38
Percent Deviation | 0 | 0.86957 | -3.0172 | 0 | -5.2174 | 2.3529 | -2.1505 | 0.50761 | -2.0101 | -0.75949 | 0
Laboratory Replicates
Masterson.UT.20240725.005: FD75 | 22.1 | 2.29 | 29.1 | 0.735 | 51.1 | 14.4 | 4.21 | 0.0192 | 4010 | 33.4 | 7.27
REP Masterson.UT.20240725.005: REP FD75 | 22 | 2.22 | 30.2 | 0.737 | 49.3 | 14.3 | 4.17 | 0.0192 | 3870 | 32.9 | 7.14
Percent Deviation | -0.45455 | -3.1532 | 3.6424 | 0.27137 | -3.6511 | -0.6993 | -0.95923 | 0 | -3.6176 | -1.5198 | -1.8207
Reference Standards (mg L-1)
| B | Ba | Ca | Fe | K | Li | Mg | Mn | Na | Si | Sr
Certified Seawater (diluted)
Actual Concentration | 5 | 0.05 | 400 | 0.02 | 380 | 0.1 | 1250 | 0.01 | 10500 | 4 | 12
Measured Concentration | 4.93 | <MDL | 391 | <MDL | 406 | <MDL | >STNDS | <MDL | 11100 | >STNDS | 12
Percent Deviation | B: -1.4199; Ca: -2.3018; K: 6.4039; Na: 5.4054; Sr: 0
M238
Actual Concentration | 0.15 | N/A | 40.9 | N/A | 5.99 | N/A | 10.1 | N/A | 22.3 | 6.45 | 0.0608
Measured Concentration | 0.137 |  | 30.7 |  | 6.05 |  | 10.1 |  | 23.3 | 5.83 | 0.0593
Percent Deviation | -9.4891 |  | -33.2248 |  | 0.99174 |  | 0 |  | 4.2918 | -10.6346 | -2.5295
M242
Actual Concentration | 0.0684 | N/A | 21.1 | N/A | 2.23 | N/A | 6.17 | N/A | 13.9 | 6.45 | 0.128
Measured Concentration | 0.0626 |  | 21.1 |  | 2.48 |  | 6.58 |  | 14.7 | 6.42 | 0.136
Percent Deviation | -9.2652 |  | 0 |  | 10.0806 |  | 6.231 |  | 5.4422 | -0.46729 | 5.8824
T241
Actual Concentration | 0.00867 | 0.054 | 18.5 | 0.232 | 1.15 | 0.012 | 6.76 | 0.0176 | 16 | 2.89 | 0.164
Measured Concentration | <PQL | 0.0544 | 18.4 | 0.237 | 1.22 | 0.0127 | 7.05 | 0.0173 | 16.4 | 2.95 | 0.17
Percent Deviation |  | 0.73529 | -0.54348 | 2.1097 | 5.7377 | 5.5118 | 4.1135 | -1.7341 | 2.439 | 2.0339 | 3.5294
T249
Actual Concentration | 0.211 | 0.11 | 83.4 | 0.0795 | 5.32 | 0.0528 | 19.8 | 0.226 | 69.7 | 6.5 | 0.404
Measured Concentration | 0.203 | 0.113 | >STND | 0.077 | 5.47 | 0.0498 | 20.5 | 0.228 | 74.4 | 6.41 | 0.408
Percent Deviation | -3.9409 | 2.6549 |  | -3.2468 | 2.7422 | -6.0241 | 3.4146 | 0.87719 | 6.3172 | -1.4041 | 0.98039
Matrix Spikes (mg L-1)
| B | Ba | Ca | Fe | K | Li | Mg | Mn | Na | Si | Sr
MS: Masterson.UT.20240725.001: FD43 | 1.84 | 0.254 | 7.93 | 0.144 | 6.19 | 0.502 | 1.46 | 0.098 | 338 | 7.16 | 2.08
1LD: Masterson.UT.20240725.001: FD43 | 1.59 | 0.161 | 3.26 | 0.046 | 1.61 | 0.239 | 0.522 | 0.003 | 280 | 5.54 | 0.62
Spike Added | 0.25 | 0.1 | 5 | 0.1 | 4.5 | 0.25 | 1 | 0.1 | 45 | 1.5 | 1.5
Percent Recovery | 100 | 93 | 93.4 | 98 | 101.7778 | 105.2 | 93.8 | 95 | 128.8889 | 108 | 97.3333
1. Corresponds to 90% (volume basis) measured values post dilution
Notes
A. Yellow highlight indicates value outside of acceptance criteria defined below.
QA-QC Acceptance Criteria: | B: Notes
1. Slope precision ≤2% at the 2s level. | B: N/A
2. Blanks <10% of PQL. | B: Potassium for blank #4 exceeds acceptance criteria. Value still below practical quantification limit (PQL) and reporting limit.
3. ICV values within ±10% of true value. | B: N/A
4. CCV values within ±10% of true value. | B: N/A
5. Lab duplicates have no acceptance criteria. | B: N/A
6. Lab replicate RPD <±20% | B: N/A
7. Reference standards within ±15% of true value. | B: One reference standard for Calcium exceeds acceptance criteria. Value is underestimated. Other reference standards and QC samples are within acceptance criteria.
8. Matrix spike recovery within 70% - 130%. | B: N/A
B. The green highlight indicates the concentrations are less than the PQL, the lowest calibration standard and the blue highlight indicates the concentrations are greater than the highest calibration standard.
Element | B: Calibration Range (mg L-1)
B | B: 0.05 - 2.5
Ba | B: 0.02 - 1
Ca | B: 1.0 - 50
Fe | B: 0.02 - 1
K | B: 0.9 - 45
Li | B: 0.05 -2.5
Mg | B: 0.2 - 10
Mn | B: 0.02 - 1
Na | B: 9 - 450
Si | B: 0.3 -15
SI | B: 0.3 -15

## Data Summary
BRInE_ID | Submitter_ID | B | ± B Error | Ba | ± Ba Error | Ca | ± Ca Error | Fe | ± Fe Error | K | ± K Error | Li | ± Li Error | Mg | ± Mg Error | Mn | ± Mn Error | Na | ± Na Error | Si | ± Si Error | Sr | ± Sr Error
|  | mg/L | mg/L | mg/L | mg/L | mg/L | mg/L | mg/L | mg/L | mg/L | mg/L | mg/L | mg/L | mg/L | mg/L | mg/L | mg/L | mg/L | mg/L | mg/L | mg/L | mg/L | mg/L
Masterson.UT.20240725.001 | FD43 | 11.3 | 0.6 | 1.14 | 0.07 | 23.9 | 1.7 | 0.328 | 0.014 | 12.1 | 1.1 | 1.66 | 0.07 | 3.8 | 0.11 | <0.125 | <PQL | 2030 | 150 | 39.8 | 1.3 | 4.38 | 0.2
Masterson.UT.20240725.002 | FD50 | 15 | 0.3 | 0.815 | 0.073 | 35.2 | 1.5 | 0.354 | 0.021 | 21.1 | 1.5 | 6.42 | 0.45 | 5.04 | 0.17 | <0.167 | <PQL | 2700 | 200 | 37.1 | 0.9 | 5.93 | 0.12
Masterson.UT.20240725.003 | FD51 | 15.3 | 0.4 | 0.244 | 0.017 | 26.8 | 1.2 | 1.34 | 0.05 | 23.8 | 2.6 | 6.66 | 0.28 | 4.23 | 0.15 | <0.143 | <PQL | 2460 | 210 | 39.5 | 0.9 | 4.63 | 0.09
Masterson.UT.20240725.004 | FD68 | 18.6 | 0.6 | 0.323 | 0.026 | 20.5 | 1.2 | 0.229 | 0.013 | 40.8 | 2.9 | 12.9 | 0.7 | 3.53 | 0.16 | <0.200 | <PQL | 3050 | 230 | 38.3 | 0.8 | 5.26 | 0.41
Masterson.UT.20240725.005 | FD75 | 22.1 | 0.8 | 2.29 | 0.14 | 29.1 | 1.2 | 0.735 | 0.031 | 51.1 | 3.2 | 14.4 | 0.6 | 4.21 | 0.12 | <0.022 | <MDL | 4010 | 290 | 33.4 | 1.3 | 7.27 | 0.45
Masterson.UT.20240725.006 | FD70 | 16.5 | 0.3 | 0.907 | 0.075 | 36.5 | 2.1 | 0.533 | 0.024 | 32.8 | 2.2 | 9.19 | 0.84 | 5.16 | 0.14 | <0.200 | <PQL | 3210 | 240 | 35.5 | 1.9 | 7.01 | 0.34
Masterson.UT.20240725.007 | FD86 | 40.2 | 1.3 | 5.22 | 0.43 | 50.3 | 2.5 | 1.48 | 0.09 | 86.8 | 9.6 | 19.4 | 0.4 | 10.1 | 0.6 | <0.050 | <MDL | 7090 | 500 | 26.2 | 0.7 | 26.2 | 1.2
Masterson.UT.20240725.008 | FD88 | 37.3 | 1.6 | 5.13 | 0.23 | 60.2 | 2.9 | 1.15 | 0.05 | 102 | 13 | 20.8 | 0.7 | 9 | 0.26 | <0.500 | <PQL | 6880 | 500 | 30.2 | 1.1 | 26.7 | 1.4
Flags
N/A
Notes
Concentrations below the method detection limit (MDL), below the lowest calibration standard (PQL), or above the highest standard (Standard) are reported as the concentration of the MDL, PQL, or high standard with the sample's dilution factor.
Error limits correspond to the ± 2𝛔 precision uncertainties. Concentrations below MDL, below PQL, or greater than the highest standard are reported as <MDL, <PQL, or >Stnd.
A. QA-QC and Sample Considerations
(i) Potassium for blank #4 exceeds acceptance criteria. Value still below practical quantification limit (PQL) and reporting limit.
(ii) One reference standard for Calcium exceeds acceptance criteria. Value is underestimated. Other reference standards and QC samples are within acceptance criteria.
B. Element concentrations based on:
Element | Line (nm)
B | 249.773
Ba | 455.403
Ca | 315.887
Fe | 259.94
Li | 670.791
Mg | 279.079
Mn | 257.61
K | 766.49
Na | 588.995
Si | 251.611
Sr | 407.771
C. Calibration standards range and MDL concentrations
Element | MDL (mg L-1) | Calibration Range (mg L-1)
B | <0.001 | 0.05 - 2.5
Ba | <0.001 | 0.02 - 1
Ca | <0.04 | 1.0 - 50
Fe | <0.0007 | 0.02 - 1
K | <0.060 | 0.9 - 45
Li | <0.0007 | 0.05 -2.5
Mg | <0.005 | 0.2 - 10
Mn | <0.0002 | 0.02 - 1
Na | <0.07 | 9 - 450
Si | <0.008 | 0.3 -15
Sr | <0.0001 | 0.3 -15
D. Sample dilution factors. Stated concentrations already reflect dilution correction.
Brine ID | Submitter ID | Dilution Factor
Masterson.UT.20240725.001 | FD43 | 6.3
Masterson.UT.20240725.002 | FD50 | 8.3
Masterson.UT.20240725.003 | FD51 | 7.1
Masterson.UT.20240725.004 | FD68 | 10
Masterson.UT.20240725.005 | FD75 | 11.1
Masterson.UT.20240725.006 | FD70 | 10
Masterson.UT.20240725.007 | FD86 | 25
Masterson.UT.20240725.008 | FD88 | 25

## Traceability
ERP-BRInE-SOP 13
SOP Version | 02
BRInE Sample ID | Masterson.UT.20240725.001-008
Date of Analysis | August 14th, 2024
Analyst | A. Herzberg
Reviewer | J. Chenault
Method Conditions
RF Power (W) | 1250
Pump Speed (rpm) | 31
Ar Plasma Gas Flow (l/min) | 14.5
Ar Aux Gas Flow (l/min) | 0.6
Ar Sheath Gas Flow (l/min) | 0.3
N2 purge flow (l/min) | 2
Nebulizer Flow (ml/min) | 0.9
Nebulizer Pressure (psi) | 2.4
Notes
1. Potassium for blank #4 exceeds acceptance criteria. Value still below practical quantification limit (PQL) and reporting limit.
2. One reference standard for Calcium exceeds acceptance criteria. Value is underestimated. Other reference standards and QC samples are within acceptance criteria.